## Slide 1: The Camera Challenge
Gavin Dalton
RALSpace

## Slide 2
Big: 2.5m
Heavy: 2.5T
Long gestation: >2yrs
Long lifetime: >25years
Difficult to handle
Difficult to transport
Requires lots of consumables
Produces vast quantities of output
Difficult to control
Difficult thermal control
Precision optics
Coplanar detectors
Cryogenic temperatures

## Slide 4: Timeline
System Design Phase (post CoDR): Apr 02
PDR held December 02
FDR held January 04
Camera fully built May 06
  Moved from assembly area
Delivery to Paranal Jan 07
On-sky mid-2008
  Commissioning of camera + telescope through mid-2009

## Slide 5: Basic Parameters
Field of View 1.67 degrees @ f/3.3 (350mm focal plane)
16 2048x2048 Raytheon VIRGO HgCdTe Arrays
(840µm-2.5µm sensitive)
YJHKs baseline filter set (one filter per detector)
3 vacant slots in wheel (Z filter to take one)
4 2048x2048 E2V CCDs for wavefront sensing
2 2048x2048 E2V FT CCDs for autoguiding
7 nested ‘dichroic’ baffles for stray light control

## Slide 6: Optical System
Telescope and camera design optimised as a single system
Full system delivers 0.5” images (50% EED). Image distortion < 2% at extreme edge of field
1m diameter

## Slide 7: Challenges
Cryostat
~1m vacuum window
Focal plane co-planarity (±25μm)
Detector procurement (ITAR)
Wavefront sensing (totally new)
Stray light baffles
Testing (optical system – 7000nm astigmatism)
Handling (in the lab and to/at the telescope)

## Slide 8: Cryostat
5 sections, 15m of o-rings
Limited possibilities for suppliers
Substantial CAD assistance needed from RAL
Supplier staffing issues (non-work related injuries)
Heat shields (use of MLI rejected)
The ‘shrinkage’ problem

## Slide 10: Handling Issues

## Slide 14: Field Corrector

## Slide 15: Baffle system
In Camera: 1 flat, 7 nested elliptical dichroic baffles, reject out-of-band thermal through dewar window.
On-telescope: double reflective baffle around M2 restricts camera view past secondary, obstruction 1.4m.

## Slide 16: Baffle Assembly

## Slide 18: Wavefront Sensors/Autoguiders

## Slide 19: Filter Wheel
Each tray has 160 separate components…

## Slide 21: Cryostat Window
0.8MJ of stored potential energy due to pressure on the window!

## Slide 22: Focal Plane Layout -> Images

## Slide 24: Focal Plane Assembly Details
Detector co-planarity: all pixels within ±25µm (Thanks Raytheon!)
ATC testing: Detector output not quite as expected… 1024 new Op Amps…

## Slide 25: The scattered light puzzle…
Dark image

## Slide 26: The scattered light puzzle…
Dark image

## Slide 27: The scattered light puzzle…
Key to explanation came from chips on the edge of one detector…
-a side effect of 800μm thick substrates… (should have been 20μm)

## Slide 28: Laboratory image tests
Two-mirror + corrector injection source to simulate telescope aberrations. Align to generate a spot in the focal plane, then measure the source configuration with Shack-Hartmann test camera.
Predicted:measured Z6,Z7, Z11=(7025,-954, 729):(-6957,944,-742nm
Cross-check Z6 measurements against through-focus (±2mm ) spot images:
Error in camera optics < 130nm (111.5mm field) or < 240nm (113.5mm field)
Notes: +300nm arises because the actual field positions used are 2.5mm or 4.5mm further off-axis than modelled.
-1160nm effect comes from an independent measurement of a systematic underestimate of the results returned by the SH system.

## Slide 29: Getting it there

## Slide 30
Shipped intact and at vacuum…
Cross-channel ferry (roughest night of the year) then 3 days by 747 (just fits). Finally by road from Santiago
Arrived at Paranal and unpacked (still at vacuum).
…not everything arrived in such good shape (3/5 He compressors damaged in transit)

## Slide 31
IR Camera in ground floor of VISTA Enclosure

## Slide 34
First Light
Early July 2008
M1/M2 setup taken from nominal test-camera data
AGs not set up
WFS -> M2 not set up
WFS-> M1not set up
WFS returning plausible Zernike values
  Apply corrections by hand, can estimate correct focus to ~200nm in Z4 by eye (WFS does a lot better).
After about 5 hours of this we stopped to take a picture…
(6x30s pawprints in J, dithered ‘by hand’)

## Slide 35: Some Images
35° elevation, 0.85” images over most of the field

## Slide 36: Wavefront Sensors
Pair of pre/post-focal images
(LOW order)
Analysis fit to the image-pair
Residuals

## Slide 37: Wavefront Sensors
High Order
25 Zernike terms fitted
(Spherical terms not applied, but no significant level detected)
M2 trefoil (~800nm) seen with test S-H camera recovered well
Can sometimes fail to converge…

## Slide 38
Beam-splitter cubes in-between filter positions can be deployed to 8 pre-determined locations for high-order curvature sensing measurements.

## Slide 39: Calibrating M1/M2
30/8/08
0.80”
0.76”
0.79”
0.82”
VLT DIMM
recording 0.8”-1.0”

## Slide 40: Residual Camera Astigmatism
Rotational variation of M1 astig corrections suggest a small camera term, consistent with the laboratory measurements.
-Add a term to the active optics controls to apply this as f(rotator) in open loop.

## Slide 41: M2 tilt (field dep. astig)
Always appeared to see worse images in one, but not always the same, corner… Z5 and Z6 measurements initially confusing.
Model as two-component rotating term, and get better data…

## Slide 42: M2 tilt (field dep. astig)
2 component model with a true field corrector mis-alignment term and a ‘phantom’ 2θ term from the gradient of the M1 trefoil correction. –Apply one term open-loop and ignore the other.

## Slide 43: Focus Gradient (HOWFS)
October 2008
Difference between focus measured at corner- and near-axis-HOWFS locations (nm of Z4

## Slide 44
Difference between extreme and near-axis positions for each rotator setting

## Slide 45: Focus Gradient (HOWFS)
March 2009
Same plot after lateral M1 position adjustment and recalibration of M2 positioning
Further repositioning of M1 in May 2009 corrected this residual

## Slide 46: Final image quality from SV data
Formal image quality 0.45” (50%EED)

## Slide 47: VISTA Performance
Throughput:
Z 51%
Y 45% (Filter)
J 53%
H 66%
Ks 63%
Total system…
telescope+camera+detectors
Array striping: controller effect, removed as identical in four arrays/IRACE sub unit
Channel cross-talk: none
Persistence: 1st Mag star residual disapppears completely in < 1min

## Slide 48: Detector FlatFields
Lower left corner

## Slide 49: Detector FlatFields
Lower right corner

## Slide 50: Detector FlatFields
Upper left corner

## Slide 51: Detector FlatFields
Upper Right corner